BronchoTrack_V3 — Progress Update

A summary of recent development, prepared for supervisor review

| Prepared by: Aveesha Nishendra | Date: September 8, 2026 |
| --- | --- |
| Project: BronchoTrack_V3 — bronchoscope airway localization |  |

This report summarizes progress on the bronchoscope tracking project since the last update. It covers a limitation identified in the existing system, an initial redesign that was implemented, evaluated, and ultimately not adopted, the solution that was implemented in its place, and the results of testing that solution on real patient video. A technical appendix at the end of this document gives the exact algorithmic detail (pseudocode for every file) behind the redesign described in the Limitation Identified section; this main body is intended as an accessible overview of the same work.

As background: the objective of this line of work is a system that can observe bronchoscope video in real time and determine, automatically, which airway branch the camera is currently positioned in, using the video feed together with a three-dimensional model of the patient's airway tree built in advance from their CT scan. No additional tracking hardware or manual annotation during the procedure is required. This report covers one complete development cycle on that system: a limitation that was identified, a first attempted fix that was set aside, and the fix that was ultimately adopted.

Project Overview

The system's objective is to determine, from live bronchoscope video alone, which airway branch the camera is currently viewing — the trachea, the left or right main bronchus, or one of their sub-branches. In practical terms, this would allow a bronchoscope to indicate the operator's current position on a map of the patient's own airways, without requiring separate tracking hardware.

This is achieved by combining two sources of information. The first is a pre-operative three-dimensional model of the patient's airway tree, built from their CT scan, which provides the known shape and branching structure in advance. The second is the live video itself: the system detects the openings in the airway wall that lead to each branch, tracks them consistently from frame to frame as the camera moves, and compares what is currently visible against what the three-dimensional model predicts should be visible from the camera's estimated position. When the two are in agreement, the system assigns a label — for example, identifying a given opening as the right upper lobe branch — and this becomes the basis for reporting the scope's location.

This work is based on a published method (BronchoTrack). The implementation maintains a strict, faithful reproduction of that method as a separate baseline, with a small number of additional, clearly documented improvements layered on top where the published method is underspecified or where practical issues were identified during development. Keeping these separate makes it possible to attribute any given result to the published method itself versus an addition made during this project.

The broader motivation is worth stating explicitly: bronchoscopists currently navigate largely by memory and visual landmarks, and disorientation within a small, branching airway is a recognized challenge, particularly for less experienced operators or unusually complex anatomy. A reliable, camera-only localization system represents a step toward tools that could assist with this in real time, without introducing additional imaging hardware into the procedure room.

System Overview

For each frame of video, the system performs four steps in sequence: it detects the airway openings visible in that frame; a tracking stage maintains each opening's identity consistently from one frame to the next, so that a given opening is not mistaken for a new one as the camera moves; an association stage determines which anatomical branch each currently visible opening most likely corresponds to, using the three-dimensional model as a reference; and finally, once openings are labeled, the system takes a vote among everything currently visible to determine the scope's overall location. The diagram below illustrates this sequence.

Figure 1. How a single video frame moves through the system, from input to output.

Two aspects of this design are worth noting. First, the detection stage produces a full outline of each opening rather than a simple bounding box; the size and shape of an opening, not merely its approximate position, provides useful evidence for identifying which branch it is, and this information would be lost with a box-only approach. Second, the system includes an independent verification step: after an opening is matched to a branch by position, its apparent size at that distance is separately checked against what the three-dimensional model predicts. A label is only treated as fully confirmed once both checks are in agreement — a stricter standard than it may initially appear, and one that becomes relevant again in the results section below.

Limitation Identified

The system was already performing reasonably well, but had a specific limitation: it only labels a new opening by comparing it against an opening it is already confident about, referred to as an “anchor.” This is a reasonable approach under normal conditions. However, if the video encounters a difficult stretch — motion blur, a brief loss of view, or a period in which the detector fails to identify anything — and every anchor is lost simultaneously, the system had no mechanism to recover. It would stop reporting a location for the remainder of the video, having lost anything to compare new detections against.

This was not a theoretical concern; it was confirmed on a real patient recording. Across a 1,084-frame video, the system lost its last anchor roughly a fifth of the way through and reported no location for approximately 82% of the remaining video. This is clearly insufficient for a system intended to track a scope's position throughout an entire procedure.

Initial Approach: A Redesign That Was Not Adopted

The first approach taken was a more substantial redesign. Rather than relying on a persistent anchor at all, this version combined three independent signals every frame: the rate at which a visible opening's apparent size was increasing, used as an indicator that the scope was approaching it; a position-based match against the three-dimensional model; and a shape-based match using the complete detected outline of each opening rather than a single reference point. These three signals were combined, with smoothing over time, into a single running confidence score per branch. The intent was that the system would never depend on any single tracked object persisting from one frame to the next, which would eliminate the failure mode described above by construction.

This approach was fully implemented, covered by an automated test suite, and evaluated on the same real video used to identify the original problem. Upon direct comparison with the existing, simpler system, however, it performed worse overall — the added complexity did not translate into a corresponding improvement in reliability. Based on this evaluation, the decision was made to set this approach aside and return to the original system, addressing the identified limitation with a smaller, more targeted change instead.

This attempt is included in this report rather than omitted, as it represents a meaningful finding in its own right: it confirmed that the existing anchor-based approach was fundamentally sound and required a bounded correction for one specific failure mode, rather than a complete redesign. That is a useful conclusion even though this particular implementation was not adopted as the primary system. It remains in the codebase, fully tested and functional, should it prove useful as a reference in future work.

Solution Implemented

The solution that was adopted retains the original anchor-based approach but adds a recovery mechanism. The system now retains a snapshot of the most recently confirmed opening, even after that opening is no longer visible. If every anchor is lost, the system attempts recovery in two stages before giving up:

- First, the system checks whether one of the currently unidentified openings is likely the same opening reappearing — that is, the scope's own current position, returning after a brief interruption — based on how closely it overlaps with the opening's last known position. If the overlap is sufficient, the opening is relabeled directly, without requiring a fresh match.

- If no sufficiently close match is found, the system falls back to using the last known position as a reference point, allowing newly visible branches nearby to still be identified against the three-dimensional model, even though the original opening itself is no longer present.

In either case, once an opening is successfully re-identified, the system resumes normal operation immediately; this is not a one-time reset, and the mechanism can recur as many times as needed over the course of a video. A safeguard is also included: if an excessive amount of time has elapsed since an anchor was last confidently observed (approximately three seconds), the system treats the retained snapshot as unreliable and withholds recovery until a new anchor is established, rather than acting on outdated information.

Results

The solution was evaluated on the same real, 1,084-frame patient video used to originally identify the limitation, comparing a run with recovery disabled against one with recovery enabled, under otherwise identical conditions:

| Metric | Before | After |
| --- | --- | --- |
| Frames with a location reported | 15.3% | 39.9% |

Table 1. Same 1,084-frame video and settings, before and after the recovery mechanism.

The proportion of the video with a reported location increased by a factor of approximately 2.6, confirming that the system no longer permanently stops reporting a location following this type of interruption.

One important caveat should be noted. A separate, stricter check governs whether a label is confident enough to be displayed on the output video, comparing the label against the three-dimensional model's expected geometry rather than tracking alone. This stricter metric remained essentially unchanged (14.9% before, 14.4% after). This is not an inconsistency: the recovery mechanism is functioning exactly as intended, restoring a label after an interruption, while this separate confidence check addresses a different question — how strictly the system should trust a given match — and is unaffected by the recovery mechanism itself. Whether this threshold should be relaxed specifically for labels restored through recovery, as opposed to those matched under normal conditions, is a design question worth discussing further.

Visualization Improvement

A secondary improvement was also made to the output video. Previously, each identified opening was represented by a small dot at its center. The output now displays the actual detected shape as a filled, semi-transparent outline, which makes it considerably easier to visually verify what the system is detecting, rather than relying on a single point. This change also surfaced a minor, pre-existing artifact: an opening located near the edge of the frame is occasionally detected with an irregular, oversized outline, because the detection boundary follows the edge of the frame rather than stopping cleanly at the true edge of the opening. This is a property of the raw detection output rather than a tracking defect, and has been noted as a candidate for a filtering step if it proves distracting in practice.

Testing and Validation

To verify the results described above, I ran the pipeline myself, on my own machine, against the same patient video I had originally provided — comparing its behavior with the recovery mechanism disabled against its behavior with the recovery mechanism enabled. This is the direct source of the before-and-after results reported in the Results section: both runs were produced under my own testing, using identical settings apart from the recovery mechanism itself, so the comparison reflects a genuine before-and-after on real video rather than a simulated or synthetic case.

Current Status and Next Steps

The recommended, actively maintained version of the system now includes the recovery mechanism and the visualization improvement described above. The earlier three-signal redesign remains in the codebase for reference but is not the direction being pursued going forward.

Two open questions would benefit from further discussion: whether the stricter confidence check should apply a more lenient standard to a label restored through the recovery mechanism, as opposed to one matched under normal conditions; and whether the frame-edge detection artifact described above warrants a dedicated fix before the next round of testing.

The exact algorithmic detail for the redesign discussed above — pseudocode for every file in that implementation — is included as a technical appendix at the end of this document. Demonstration videos are available and can be provided on request.

Sincerely,

Aveesha Nishendra

Technical Appendix: Redesign Implementation Reference

This appendix gives the exact algorithmic detail behind the three-signal redesign discussed in the Limitation Identified section above (the diameter-growth motion model, position-based bearing matching, and whole-mask ratio method, combined through a Kalman filter). It is not the recommended, actively maintained system — see Current Status and Next Steps — but is included here as a complete reference, since it remains in the codebase and represents a meaningful part of this development cycle. Each entry below gives one file's role in a line, followed by a functional walkthrough of its classes and functions in the order they execute.

bronchotrack/scalar_kalman.py

Shared 2-state constant-velocity Kalman filter ([value, rate]), reused by both paper_exact (roll angle, diameter:distance smoothing) and fusion (diameter growth rate, per-label confidence smoothing).

function build_matrices(dt): F = [[1, dt], [0, 1]] # state transition: value += rate*dt H = [[1, 0]] # we only observe "value", not "rate" return F, H class ScalarKalmanFilter: state: [value, rate], covariance P, process_noise Q, measurement_noise R function predict(dt): F, _ = build_matrices(dt) state = F . state P = F . P . F_transpose + Q return self function update(measurement): F, H = build_matrices(dt=1) y = measurement - H . state # innovation S = H . P . H_transpose + R K = P . H_transpose . inverse(S) # Kalman gain state = state + K . y P = (I - K . H) . P return self property value: state[0] property rate: state[1]

bronchotrack/fusion/motion_model.py

Diameter-growth motion model: tracks each tracklet's apparent diameter (Kalman-smoothed) and derives a scale-free time-to-contact. Cross-checked (never calibrated) against the airway graph's known branch lengths.

class ApproachEstimate: fields: track_id, diameter, growth_rate, approaching (bool), tau_frames (nullable) function detection_diameter(tracklet): return tracklet's box height # proxy for apparent lumen size class ApproachMotionModel: filters: map of track_id -> ScalarKalmanFilter # smooths diameter + its rate of change function update(tracklets, frame_idx): estimates = {} for t in tracklets: diam = detection_diameter(t) filt = filters.get(t.track_id) or new ScalarKalmanFilter(seeded at diam) filt.predict(dt=1); filt.update(diam) filters[t.track_id] = filt growth_rate = filt.rate approaching = growth_rate > small_epsilon # only "approaching" if genuinely growing tau_frames = filt.value / growth_rate if approaching else None # tau = diameter / d(diameter)/dt -- scale-free time-to-contact estimates[t.track_id] = ApproachEstimate(t.track_id, filt.value, growth_rate, approaching, tau_frames) return estimates function tau_frames_for(track_id): return estimates.get(track_id).tau_frames if present else None function prune(live_track_ids): remove any entry in filters whose track_id is not in live_track_ids class TransitRecord: fields: label, frames_spent, arc_length_mm class TransitSanityChecker: plausible_range: (low, high) # mm per frame bounds history: list of TransitRecord function record_transit(label, frames_spent, arc_length_mm): append TransitRecord to history property median_speed_mm_per_frame: return median(arc_length_mm / frames_spent for each record), or None if empty function check(tau_frames, candidate_arc_length_mm): if history is empty: return "unknown" implied_speed = candidate_arc_length_mm / tau_frames ratio = implied_speed / median_speed_mm_per_frame return "plausible" if plausible_range.low <= ratio <= plausible_range.high else "implausible" # empirical, self-referential check, not a calibration

bronchotrack/fusion/bronchotrack_id.py

'The bronchotrack version': point-based bearing matching using each lumen's segmentation center, re-derived fresh rather than importing paper_exact.association's implementation.

function detection_point(tracklet): return polygon_opposing_center(tracklet.mask) if tracklet has a mask else tracklet's box center function bearing_cost_matrix(graph, ref_label, candidate_labels, detections, reference_point): if candidate_labels is empty: return None projected = graph.project_children_2d(ref_label) # 2D unit bearing per candidate child if any candidate label is missing a projection: return None cost = matrix[len(candidate_labels)][len(detections)] for i, label in candidate_labels: for j, det in detections: observed_bearing = normalize(detection_point(det) - reference_point) cost[i][j] = 1 - cosine_similarity(observed_bearing, projected[label]) return cost # 0 = perfect bearing match, up to 2 = opposite direction

bronchotrack/fusion/ratio_id.py

'The ratio method': genuinely whole-mask, not point-based. Size is the mask's own area-equivalent diameter; distance is the nearest boundary-to-boundary gap between two full mask polygons.

function mask_boundary_gap(poly_a, poly_b): if either polygon is None: return None return min distance over all point pairs (p in poly_a, q in poly_b) # nearest boundary-to-boundary gap, an all-pairs search over both boundaries function circle_boundary_gap(center_a, radius_a, center_b, radius_b): d = distance(center_a, center_b) - radius_a - radius_b return max(d, MIN_DIST) # clamp so overlapping circles never go to zero/negative function reference_opening_radius(graph, ref_label): return graph node ref_label's known physical radius, or None if graph has no radius data function ratio_cost_matrix(graph, ref_label, candidate_labels, reference_detection, detections, virtual_advance_mm): ref_radius = reference_opening_radius(graph, ref_label) if ref_radius is None or detections is empty or reference_detection.mask is None: return None ref_equiv_diameter = mask_equivalent_diameter(reference_detection.mask) # area-based (shoelace integral over the mask boundary), not longest-diameter px_per_mm = ref_equiv_diameter / (2 * ref_radius) cost = matrix[len(candidate_labels)][len(detections)] for i, label in candidate_labels: child_radius = graph.child_apparent_equivalent_diameter_at_distance( ref_label, label, virtual_advance_mm) / 2 for j, det in detections: if det.mask is None: cost[i][j] = large_penalty; continue gap_px = mask_boundary_gap(reference_detection.mask, det.mask) gap_mm = gap_px / px_per_mm expected_gap_mm = circle_boundary_gap(using ref_radius and child_radius) cost[i][j] = abs(gap_mm - expected_gap_mm) # how well this detection's real-world spacing matches the graph's expectation return cost

bronchotrack/fusion/kalman_fusion.py

Blends the bearing and ratio cost matrices into one Hungarian assignment per frame, nudges the result by the motion model, and smooths each candidate's confidence with a per-label Kalman filter.

class CandidateConfidence: fields: label, fused_confidence, matched_track_id, raw_confidence class IdentityFusion: bearing_weight, ratio_weight, motion_gate_strength per_label_filters: map of label -> ScalarKalmanFilter # smooths confidence over time function update(labels, detections, bearing_cost, ratio_cost, approach_estimates): blended_cost = blend(bearing_cost, ratio_cost, bearing_weight, ratio_weight) # if one cue is missing (None), fall back to using only the other, unweighted row_idx, col_idx = hungarian_assignment(blended_cost) # one detection per candidate label, minimizing total cost results = {} for each (label, detection) matched by the assignment: raw_confidence = 1 - blended_cost[label][detection] gated_confidence = apply_motion_gate(raw_confidence, approach_estimates.get(detection.track_id)) filt = per_label_filters.get(label) or new ScalarKalmanFilter filt.predict(dt=1); filt.update(gated_confidence) fused = filt.value results[label] = CandidateConfidence(label, fused, detection.track_id, raw_confidence) return results function apply_motion_gate(confidence, estimate): if estimate is None: return confidence return confidence + motion_gate_strength if estimate.approaching else confidence - motion_gate_strength

bronchotrack/fusion/association.py

End-to-end continuous re-evaluation: re-runs the fused match every frame against whatever is currently visible, and commits a branch transition only after a consecutive-frame confidence streak.

class AssociationResult: fields: location, committed_this_frame, transition_sanity, candidates function detection_size(tracklet): same as motion_model's diameter proxy (box height) class FusionAssociation: graph, commit_threshold, commit_frames current_location, has_committed_once = False commit_streak: map of label -> consecutive-frame count motion_model = ApproachMotionModel() sanity_checker = TransitSanityChecker() fusion_by_location: map of location -> IdentityFusion # one fusion instance per branch point eligibility_fn: always returns True (no tracker dropout in this design) function process_frame(tracklets, frame_idx): approach_estimates = motion_model.update(tracklets, frame_idx) motion_model.prune(live track ids) if not has_committed_once: return try_bootstrap(tracklets, frame_idx) reference = pick_reference(tracklets) # the tracklet currently labeled as occupying current_location if reference is None: return AssociationResult(current_location, committed=False, sanity="unknown", candidates={}) children = graph's children of current_location fusion = get_fusion(current_location) # lazily created IdentityFusion per location bearing_cost = bearing_cost_matrix(graph, current_location, children, tracklets, detection_point(reference)) ratio_cost = ratio_cost_matrix(graph, current_location, children, reference, tracklets, virtual_advance_mm) candidates = fusion.update(children, tracklets, bearing_cost, ratio_cost, approach_estimates) committed, sanity_verdict = try_commit(candidates, tracklets, frame_idx) return AssociationResult(current_location, committed, sanity_verdict or "unknown", candidates) function pick_reference(tracklets): return the tracklet whose .label == current_location, if any, else None function try_bootstrap(tracklets, frame_idx): # special one-time case: at the very first bifurcation there is no # pre-existing occupant, so both visible lumens are labeled from # bearing alone in a single shot children = graph's children of the graph root match each visible tracklet to the nearest child bearing (bearing_cost_matrix only, no ratio/motion yet) label every matched tracklet accordingly current_location = label of the larger / most prominent matched tracklet has_committed_once = True return AssociationResult(current_location, committed=True, sanity="unknown", candidates={}) function try_commit(candidates, tracklets, frame_idx): for label, cand in candidates: if cand.fused_confidence >= commit_threshold: commit_streak[label] += 1 else: commit_streak[label] = 0 # reset immediately, not decayed best_label = candidate with highest streak that has reached commit_frames if no such label: return (False, None) matched_tracklet = tracklet with track_id == candidates[best_label].matched_track_id arc_length_mm = graph distance from current_location to best_label tau = motion_model.tau_frames_for(matched_tracklet.track_id) sanity = sanity_checker.check(tau, arc_length_mm) if tau is not None else "unknown" if sanity resolves and location actually changes: sanity_checker.record_transit(best_label, frames_spent=commit_streak[best_label], arc_length_mm) current_location = best_label reset commit_streak for all labels update_gallery(matched_tracklet, best_label) return (True, sanity) function update_gallery(tracklet, label): tracklet.label = label # so pick_reference finds it next frame function generation(label): return graph node's generation depth (used for HUD / logging)

bronchotrack/fusion/pipeline.py

Drives detector, tracker, and association together frame by frame, writes the overlay video and the JSON log.

class FrameResult: fields: frame_idx, location, committed_this_frame, candidates, tracklets class FusionPipeline: detector, tracker, association: FusionAssociation, graph function process_frame(frame_bgr, frame_idx): detections = detector.infer(frame_bgr) active_tracklets = tracker.update(detections, frame_idx, eligibility_fn=association.eligibility_fn) result = association.process_frame(active_tracklets, frame_idx) return frame_log_entry(frame_idx, result, active_tracklets) function run_on_video(video_path, out_video_path, out_json_path): open video reader, video writer log = [] for each frame, frame_idx in video: result = process_frame(frame, frame_idx) overlay_frame = viz.draw_overlay(frame, result, graph) write overlay_frame to out_video log.append(result) write_json_log(log, out_json_path) close reader/writer function frame_log_entry(frame_idx, result, tracklets): return dict with frame_idx, current location, committed flag, per-candidate confidences, per-tracklet box/label info function write_json_log(log, path): serialize log list to JSON file

bronchotrack/fusion/viz.py

Draws the live overlay: tracked lumens plus a HUD panel listing every currently scored candidate child branch (not fabricated screen positions, since no camera calibration exists in this pipeline).

function draw_overlay(frame_bgr, result, graph): out = copy of frame_bgr for each tracklet currently visible: draw a dot / box at its position # no full mask HUD placement -- no camera calibration exists to place # candidate branches at real screen positions draw_candidate_hud(out, result) # side panel listing every scored candidate child draw current_location label near top of frame return out function draw_candidate_hud(out, result): panel_x, panel_y = fixed HUD corner for each (label, candidate_confidence) in result.candidates, sorted by confidence descending: draw label name draw a horizontal bar sized proportional to fused_confidence draw ETA text (from tau_frames if available, else a placeholder dash) advance panel_y for next row

bronchotrack/fusion/cli.py

Command-line entry point: parses arguments, wires up the detector, tracker, association, and pipeline, then runs the video.

function build_arg_parser(): args: --video, --graph, --weights, --conf-threshold, --out-video, --out-json, --commit-threshold, --commit-frames, --show-graph (mirrors paper_exact's CLI) return parser function main(argv): args = build_arg_parser().parse_args(argv) graph = AirwayGraph.from_json(args.graph) detector = load detector from args.weights tracker = MultiLumenTracker() association = FusionAssociation(graph, commit_threshold=args.commit_threshold, commit_frames=args.commit_frames) pipeline = FusionPipeline(detector, tracker, association, graph) pipeline.run_on_video(args.video, args.out_video, args.out_json) print("Processed <N> frames... Final location: <label> (generation <g>)")

bronchotrack/fusion/__init__.py

Package entry point: exports the public classes and documents the three fused models and the deliberate scope reductions for this first version.

# Public exports: from .association import AssociationResult, FusionAssociation from .kalman_fusion import CandidateConfidence, IdentityFusion from .motion_model import ApproachEstimate, ApproachMotionModel, TransitSanityChecker from .pipeline import FrameResult, FusionPipeline # __all__ = [FusionAssociation, AssociationResult, IdentityFusion, CandidateConfidence, # ApproachMotionModel, ApproachEstimate, TransitSanityChecker, # FusionPipeline, FrameResult] # No executable logic beyond imports/exports -- this module is the package's # documented overview: # 1. motion_model.ApproachMotionModel -- diameter-growth, scale-free tau # 2. bronchotrack_id -- point-based bearing matching (segmentation center) # 3. ratio_id -- whole-mask ratio method (area + boundary-gap, not points) # kalman_fusion.IdentityFusion blends (1) and (2) into one Hungarian # assignment per frame; association.FusionAssociation re-runs this every # frame (continuous re-evaluation) and commits only after a consecutive # confidence streak. # # Deliberate scope reductions for this first version (flagged, not hidden): # - no roll-angle correction in bronchotrack_id (plain unrotated bearing) # - no multi-generation "ambiguous cohort" localization smoothing # - no tracker eligibility dropout (eligibility_fn always True)

tests/test_fusion_smoke.py

Smoke and regression tests for the fusion package: geometry primitives, motion model behavior, fusion blending, commit hysteresis, and a no-ML end-to-end run. 18 tests total.

helper _rect_mask(cx, cy, half): builds a square polygon mask centered at (cx, cy) helper _mk_tracklet(track_id, cx, cy, h, mask, label, hits): builds a minimal test Tracklet test_child_apparent_equivalent_diameter_matches_ellipse_area_formula: build a straight branch and a 60-degree-off-axis branch off a common root assert straight branch reads ~true diameter; off-axis area-based estimate is larger than the off-axis longest-diameter estimate (matches the derived cos vs sqrt(cos) foreshortening formulas exactly) test_mask_boundary_gap_and_circle_boundary_gap_exact: two rectangles separated by a known horizontal gap -> mask_boundary_gap returns that exact gap two circles at known centers/radii -> circle_boundary_gap returns the exact gap; overlapping circles clamp to MIN_DIST test_ratio_method_disambiguates_where_bearing_alone_cannot: construct two candidate branches with identical projected bearing but different radius, two detections at the identical image point but different mask size assert bearing_cost_matrix is completely flat (no discriminating power) assert ratio_cost_matrix correctly favors the size-matched pairing test_ratio_cost_matrix_falls_back_gracefully_without_radius_data: graph with no radius data -> ratio_cost_matrix returns None test_ratio_cost_matrix_handles_no_other_detections_this_frame: empty detections list -> ratio_cost_matrix returns None cleanly, no NaN / RuntimeWarning test_bearing_cost_matrix_hungarian_correct_assignment: two candidates at right-angle bearings, two detections offset to match each assert cost matrix is exactly [[0,1],[1,0]] test_motion_model_tau_matches_hand_calculation: feed linearly growing diameter for 40 frames assert estimate.approaching is True and tau_frames is approx diameter / growth_rate test_motion_model_not_approaching_when_shrinking: feed shrinking diameter -> approaching is False, tau_frames is None test_motion_model_prune_drops_stale_filters: update with one tracklet, then prune with an empty live set -> filter removed test_transit_sanity_checker_plausibility: no history -> "unknown" record two consistent transits -> median speed computed correctly a consistent candidate -> "plausible"; a wildly different one -> "implausible" test_identity_fusion_blends_cues_and_motion_gates: two candidates with identical bearing+ratio cost, but only one detection is "approaching" assert the approaching candidate's fused_confidence ends up higher test_identity_fusion_falls_back_to_single_cue_when_other_missing: bearing_cost = None -> fusion uses ratio_cost alone, unweighted test_fusion_association_root_bootstrap_labels_both_lumens_at_once: two lumens visible at the very first frame, no prior occupant assert both get labeled correctly in one frame and the larger becomes current_location test_fusion_association_commits_after_consecutive_frames_not_one_lucky_frame: run commit_frames-1 strong frames -> not yet committed run one more strong frame -> commits to the correct child test_fusion_association_streak_resets_on_interrupting_bad_frame: 4 clearing frames, 1 frame below threshold, 4 more clearing frames (commit_frames=5) assert it never commits, because the streak resets on the bad frame instead of just pausing test_fusion_association_continuous_reevaluation_survives_reference_loss: bootstrap, then a frame with zero tracklets assert no crash and current_location is unchanged test_scalar_kalman_filter_promoted_and_reexported_identically: assert bronchotrack.scalar_kalman.ScalarKalmanFilter is the same object as bronchotrack.paper_exact.scalar_kalman.ScalarKalmanFilter (re-export, not a copy) test_full_fusion_pipeline_smoke_no_ml: run tracker + FusionAssociation together on synthetic carina-region detections (no real ML) assert both main bronchi get bootstrapped and labeled correctly _run_all(): runs every test above in sequence, prints PASS/FAIL per test and a final N/M summary